| Name | Security Level |
| --- | --- |
| BELLSINGTM Hornet Series AMBA Receiver | Open |
| Version | A Total of 4 Pages |
| REV X1 |  |

BRC910C22955S

Product Data Sheet

| Prepared by: | Zoe Wang |  | Date: | 2019-05-23 |
| --- | --- | --- | --- | --- |
| Auditor: |  |  | Date: |  |
| Approver: |  |  | Date: |  |

Shenzhen Bellsing Acoustic Tech. Co., Ltd.

BRC910C22955S AMBA receiver

  - Description

Miniature balanced-armature receiver with magnetic-radiation shielding for use in BTE (Behind the Ear) and headset as a woofer.

| Name: | Hornet Series AMBA Receiver |
| --- | --- |
| Part Number: | BRC910C22955S REV X1 |

  - Construction Specification

Weight:  0.24 g

Mechanical dimension：please refer to the drawing below

*A positive voltage applied to the negative terminal (­) will result in an increase in pressure at the sound outlet.

  - Features&Benefits

- Perfect package size for BTE , headset applications

- Outstanding efficiency

- Substantial maximum output

- Magnetic-radiation shielded

- Optimized mechanical shock resistance

  - Electro-acoustic Specifications

Test condition

- Acoustic tubing: 10.0mm of 1.0mm diameter tubing

- Acoutic coupler: IEC711 coupler

- Constant voltage drive: 0.3V RMS

- Environment: 23℃, RH 50%

Acoustic Parameter I

| Test Item |  | Nominal | Tolerance | Comments |
| --- | --- | --- | --- | --- |
|  | @20Hz | 130 | +/- 3.0 |  |
| Sensitivity(dB) | @200Hz | 129 | +/- 3.0 |  |
|  | @500Hz | 125 | +/- 3.0 |  |
|  | @1000Hz | 124.5 | +/- 3.0 |  |
| Peak1 | Frequency (Hz) | 1500 | +/- 350 |  |
|  | SPL(dB) | 129 | +/- 3.0 |  |
| Peak2 | Frequency (Hz) | 4000 | +/- 500 |  |
|  | SPL(dB) | 121 | +/- 3.0 |  |
| Maximum output@Peak frequency (dB) |  | N/A | N/A | THD<10% |

Acoustic Parameter II

| Test Item |  | Nominal | Comments |
| --- | --- | --- | --- |
| THD(%) | @1/2 1st peak | <5.0 |  |
|  | @1/3 1st peak | <5.0 |  |

Electric Parameter

| Test Item |  | Nominal | Tolerance | Comments |
| --- | --- | --- | --- | --- |
| Impedance(Ω) | @500Hz | 57.5 | +/- 15% |  |
|  | @1000Hz | 134 | +/- 15% |  |
| DCR(Ω) |  | 20 | +/- 10% |  |

Nominal Frequency Response Curve

  - Package and Storage Condition

- Package: PET bubble pack, 1 pcs per bubble, 25 pcs per pack

- Storage temperature range: -40℃ ~ 60℃

- Storage humidity range: <60%